## summaryO3 18 Jun 16
Averages of total ozone:
With the exisiting Calibration, Lamp tests included
Average is over the complete Mu range
D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD
278.3974 | 294.1189 | 279.1877 | 266.253 | 275.4039
2.9935 | 18.715 | 3.7838 | -9.1508
With Langley Plot Calibration
Langley Plot Adjustments and Lamp Test results included
D083_TO3_C* | D083_TO3_D* | D083_TO3_A* | D083_TO3_CD* | D083_TO3_AD*
275.4705 | 284.433 | 277.1703 | 268.5578 | 275.3244
0.14618 | 9.1086 | 1.8459 | -6.7665

## IdD083
Dobson 083 at IZO 2016.Sep.24 Langley Plot Results
SLC083Na | -0.8 | SLC083Nc | -0.8 | [9 empty] | SLC083Nd | -0.8 | [11 empty] | Correction to
D083 LP adj. A:-0.60 C:-0.51 D:-0.70 Nad:0.10 | [12 empty] | MaxMu | 2.5 | MaxMu can be changed | [5 empty] | LPC_C_83 |  | LPC_D_83 |  | LPC_A_83 | AD | CD
[20 empty] | With out High mu values | -0.49992 |  | -0.67498 |  | -0.49117 | 0.18382 | 0.17506 | [14 empty] | Limited Mu range
[11 empty] | 37 |  |  |  | 278.3974 | 294.1189 | 279.1877 | 266.253 | 275.4039 | all Values | -0.50649 |  | -0.70031 |  | -0.60235 | 0.09796 | 0.19381 | [14 empty] | 0.17253 |  | 0.19248
D083RC | D083RD | D083RA | day | h | m | s | m | year |  |  |  | D083NC | D083ND | D083NA | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | 1/mu.C | P*C_83 | 1/Mu.D | P*D_83 | 1/Mu.A | P*A_83 | lst.C | Mu.C | AirM.C | lst.D | Mu.D | AirM.D | lst.A | Mu.A | AirM.A | lst.CD | Mu.CD | AirM.CD | lst.AD | Mu.AD | AirM.AD | 1/mu.cd | P*CD_83 | 1/mu.ad | P*AD_83
97.9 | 55.4 | 181 | 24 | 8 | 28 | 0 | 9 | 16 | C | 83 | IZO | 91.7643 | 53.6489 | 168.036 | 276.9179 | 289.6453 | 278.056 | 267.1545 | 275.0962 | 0.34153 | 0.22887 | 0.34494 | 0.10445 | 0.34836 | 0.49705 | 2016-09-24 08:28 | 2.928 | 2.9879 | 2016-09-24 08:29 | 2.899 | 2.9569 | 2016-09-24 08:30 | 2.8706 | 2.9265 | 2016-09-24 08:28 | 2.9135 | 2.9723 | 2016-09-24 08:29 | 2.8847 | 2.9416 | 0.34323 | 0.12442 | 0.34665 | 0.3926
93.9 | 53.1 | 174.5 | 24 | 8 | 32 | 0 | 9 | 16 | C | 83 | IZO | 88.0564 | 51.4978 | 161.6729 | 276.3223 | 288.9268 | 278.0525 | 266.644 | 275.2766 | 0.35517 | 0.22838 | 0.35858 | 0.10419 | 0.36197 | 0.49704 | 2016-09-24 08:32 | 2.8155 | 2.8679 | 2016-09-24 08:33 | 2.7888 | 2.8395 | 2016-09-24 08:34 | 2.7626 | 2.8117 | 2016-09-24 08:33 | 2.8021 | 2.8536 | 2016-09-24 08:34 | 2.7757 | 2.8255 | 0.35687 | 0.12419 | 0.36027 | 0.39285
90.7 | 51.1 | 168.5 | 24 | 8 | 36 | 0 | 9 | 16 | C | 83 | IZO | 85.0912 | 49.6126 | 155.8082 | 277.7214 | 289.1807 | 278.0515 | 268.9211 | 275.2139 | 0.36875 | 0.22954 | 0.37213 | 0.10428 | 0.37551 | 0.49704 | 2016-09-24 08:36 | 2.7119 | 2.7578 | 2016-09-24 08:37 | 2.6872 | 2.7318 | 2016-09-24 08:38 | 2.663 | 2.7062 | 2016-09-24 08:36 | 2.6995 | 2.7447 | 2016-09-24 08:38 | 2.6751 | 2.7189 | 0.37044 | 0.12526 | 0.37382 | 0.39276
87.2 | 49.2 | 163 | 24 | 8 | 40 | 0 | 9 | 16 | C | 83 | IZO | 81.8448 | 47.8108 | 150.4253 | 276.748 | 288.7614 | 278.1321 | 267.5114 | 275.4233 | 0.38225 | 0.22873 | 0.38561 | 0.10413 | 0.38897 | 0.49718 | 2016-09-24 08:40 | 2.6161 | 2.6567 | 2016-09-24 08:41 | 2.5933 | 2.6327 | 2016-09-24 08:42 | 2.5709 | 2.6091 | 2016-09-24 08:40 | 2.6046 | 2.6446 | 2016-09-24 08:42 | 2.5821 | 2.6208 | 0.38393 | 0.1246 | 0.38729 | 0.39305
84.2 | 47.6 | 158 | 24 | 8 | 44 | 0 | 9 | 16 | C | 83 | IZO | 79.071 | 46.2847 | 145.5621 | 276.8784 | 289.8994 | 278.4916 | 266.8557 | 275.588 | 0.39566 | 0.22884 | 0.399 | 0.10454 | 0.40233 | 0.49783 | 2016-09-24 08:44 | 2.5274 | 2.5634 | 2016-09-24 08:45 | 2.5063 | 2.5412 | 2016-09-24 08:46 | 2.4855 | 2.5194 | 2016-09-24 08:44 | 2.5168 | 2.5522 | 2016-09-24 08:46 | 2.4959 | 2.5303 | 0.39733 | 0.1243 | 0.40066 | 0.39328
81.5 | 46.1 | 153.3 | 24 | 8 | 48 | 0 | 9 | 16 | C | 83 | IZO | 76.586 | 44.8462 | 141.012 | 277.4365 | 290.7581 | 278.7688 | 267.1757 | 275.7199 | 0.40898 | 0.2293 | 0.41229 | 0.10485 | 0.4156 | 0.49832 | 2016-09-24 08:48 | 2.4451 | 2.4771 | 2016-09-24 08:49 | 2.4255 | 2.4566 | 2016-09-24 08:50 | 2.4062 | 2.4364 | 2016-09-24 08:49 | 2.4353 | 2.4668 | 2016-09-24 08:50 | 2.4158 | 2.4464 | 0.41063 | 0.12445 | 0.41394 | 0.39347
78.8 | 44.6 | 148.9 | 24 | 8 | 52 | 0 | 9 | 16 | C | 83 | IZO | 74.1026 | 43.4005 | 136.7484 | 277.1 | 290.3231 | 278.973 | 266.9108 | 276.0872 | 0.42219 | 0.22902 | 0.42548 | 0.10469 | 0.42876 | 0.49869 | 2016-09-24 08:52 | 2.3686 | 2.3971 | 2016-09-24 08:53 | 2.3503 | 2.3781 | 2016-09-24 08:54 | 2.3323 | 2.3593 | 2016-09-24 08:52 | 2.3594 | 2.3876 | 2016-09-24 08:54 | 2.3413 | 2.3686 | 0.42384 | 0.12433 | 0.42712 | 0.39399
76.3 | 43.3 | 144.5 | 24 | 8 | 56 | 0 | 9 | 16 | C | 83 | IZO | 71.7995 | 42.1418 | 132.531 | 276.816 | 290.9561 | 278.5772 | 265.9129 | 275.4327 | 0.4353 | 0.22879 | 0.43856 | 0.10492 | 0.44181 | 0.49798 | 2016-09-24 08:56 | 2.2973 | 2.3228 | 2016-09-24 08:57 | 2.2802 | 2.3051 | 2016-09-24 08:58 | 2.2634 | 2.2876 | 2016-09-24 08:56 | 2.2887 | 2.3139 | 2016-09-24 08:58 | 2.2718 | 2.2963 | 0.43693 | 0.12387 | 0.44019 | 0.39306
74.2 | 42 | 140.6 | 24 | 9 | 0 | 0 | 9 | 16 | C | 83 | IZO | 69.8642 | 40.8774 | 128.8084 | 277.6919 | 290.4554 | 278.7367 | 267.8508 | 275.7602 | 0.44829 | 0.22951 | 0.45152 | 0.10474 | 0.45475 | 0.49826 | 2016-09-24 09:00 | 2.2307 | 2.2536 | 2016-09-24 09:01 | 2.2147 | 2.2371 | 2016-09-24 09:02 | 2.199 | 2.2208 | 2016-09-24 09:00 | 2.2227 | 2.2453 | 2016-09-24 09:02 | 2.2068 | 2.2289 | 0.44991 | 0.12477 | 0.45314 | 0.39352
72 | 41 | 137.1 | 24 | 9 | 4 | 0 | 9 | 16 | C | 83 | IZO | 67.8369 | 39.9006 | 125.4736 | 277.3396 | 292.6011 | 279.2826 | 265.5619 | 275.9028 | 0.46117 | 0.22922 | 0.46437 | 0.10552 | 0.46756 | 0.49924 | 2016-09-24 09:04 | 2.1684 | 2.189 | 2016-09-24 09:05 | 2.1535 | 2.1736 | 2016-09-24 09:06 | 2.1388 | 2.1584 | 2016-09-24 09:05 | 2.1609 | 2.1813 | 2016-09-24 09:06 | 2.1461 | 2.1659 | 0.46277 | 0.12371 | 0.46597 | 0.39372
70 | 39.8 | 133.5 | 24 | 9 | 8 | 0 | 9 | 16 | C | 83 | IZO | 65.9952 | 38.7236 | 122.036 | 277.3147 | 291.2505 | 278.9844 | 266.5603 | 275.8715 | 0.47392 | 0.2292 | 0.47709 | 0.10503 | 0.48025 | 0.49871 | 2016-09-24 09:08 | 2.1101 | 2.1287 | 2016-09-24 09:09 | 2.096 | 2.1142 | 2016-09-24 09:10 | 2.0823 | 2.1 | 2016-09-24 09:08 | 2.103 | 2.1214 | 2016-09-24 09:10 | 2.0891 | 2.1071 | 0.47551 | 0.12417 | 0.47867 | 0.39368
68.1 | 38.9 | 130.2 | 24 | 9 | 12 | 0 | 9 | 16 | C | 83 | IZO | 64.2462 | 37.837 | 118.8822 | 277.1597 | 292.8638 | 278.8899 | 265.0341 | 275.3463 | 0.48654 | 0.22907 | 0.48968 | 0.10561 | 0.4928 | 0.49854 | 2016-09-24 09:12 | 2.0553 | 2.0722 | 2016-09-24 09:13 | 2.0422 | 2.0586 | 2016-09-24 09:14 | 2.0292 | 2.0453 | 2016-09-24 09:12 | 2.0487 | 2.0654 | 2016-09-24 09:14 | 2.0357 | 2.0519 | 0.48811 | 0.12346 | 0.49124 | 0.39293
66.5 | 38.1 | 127.3 | 24 | 9 | 16 | 0 | 9 | 16 | C | 83 | IZO | 62.7717 | 37.0462 | 116.1169 | 278.021 | 295.0411 | 279.3592 | 264.8746 | 275.3847 | 0.49903 | 0.22979 | 0.50213 | 0.10639 | 0.50522 | 0.49938 | 2016-09-24 09:16 | 2.0039 | 2.0192 | 2016-09-24 09:17 | 1.9915 | 2.0064 | 2016-09-24 09:18 | 1.9793 | 1.9939 | 2016-09-24 09:16 | 1.9977 | 2.0128 | 2016-09-24 09:18 | 1.9854 | 2.0001 | 0.50058 | 0.12339 | 0.50368 | 0.39298
55.9 | 32.2 | 107.9 | 24 | 9 | 48 | 0 | 9 | 16 | C | 83 | IZO | 52.9351 | 31.1541 | 97.8697 | 279.4605 | 295.0365 | 279.3142 | 267.4208 | 275.333 | 0.5935 | 0.23097 | 0.59629 | 0.10639 | 0.59906 | 0.4993 | 2016-09-24 09:48 | 1.6849 | 1.6923 | 2016-09-24 09:49 | 1.677 | 1.6843 | 2016-09-24 09:50 | 1.6693 | 1.6764 | 2016-09-24 09:48 | 1.681 | 1.6883 | 2016-09-24 09:50 | 1.6732 | 1.6803 | 0.59489 | 0.12458 | 0.59767 | 0.3929
55 | 31.5 | 106.1 | 24 | 9 | 52 | 0 | 9 | 16 | C | 83 | IZO | 52.0934 | 30.4493 | 96.1962 | 280.4205 | 292.7508 | 279.6271 | 270.8923 | 276.3028 | 0.60457 | 0.23177 | 0.60731 | 0.10557 | 0.61004 | 0.49986 | 2016-09-24 09:52 | 1.6541 | 1.6609 | 2016-09-24 09:53 | 1.6466 | 1.6533 | 2016-09-24 09:54 | 1.6392 | 1.6458 | 2016-09-24 09:52 | 1.6503 | 1.657 | 2016-09-24 09:54 | 1.6429 | 1.6495 | 0.60594 | 0.1262 | 0.60867 | 0.39429
53.7 | 31 | 104.2 | 24 | 9 | 56 | 0 | 9 | 16 | C | 83 | IZO | 50.8743 | 29.945 | 94.4251 | 278.2319 | 293.3197 | 279.2993 | 266.5684 | 275.7487 | 0.61546 | 0.22996 | 0.61815 | 0.10577 | 0.62083 | 0.49927 | 2016-09-24 09:56 | 1.6248 | 1.6311 | 2016-09-24 09:57 | 1.6177 | 1.6238 | 2016-09-24 09:58 | 1.6107 | 1.6167 | 2016-09-24 09:56 | 1.6213 | 1.6274 | 2016-09-24 09:58 | 1.6142 | 1.6203 | 0.61681 | 0.12419 | 0.6195 | 0.3935
52.8 | 30.5 | 102.4 | 24 | 10 | 0 | 0 | 9 | 16 | C | 83 | IZO | 50.0274 | 29.44 | 92.7388 | 278.431 | 293.4193 | 278.9612 | 266.8438 | 275.3003 | 0.62617 | 0.23012 | 0.62881 | 0.10581 | 0.63145 | 0.49867 | 2016-09-24 10:00 | 1.597 | 1.6028 | 2016-09-24 10:01 | 1.5903 | 1.5959 | 2016-09-24 10:02 | 1.5837 | 1.5892 | 2016-09-24 10:00 | 1.5936 | 1.5993 | 2016-09-24 10:02 | 1.587 | 1.5925 | 0.62749 | 0.12431 | 0.63013 | 0.39286
51.9 | 30.1 | 100.8 | 24 | 10 | 4 | 0 | 9 | 16 | C | 83 | IZO | 49.1783 | 29.0354 | 91.2357 | 278.2813 | 294.8478 | 278.9843 | 265.4719 | 274.9685 | 0.63669 | 0.23 | 0.63929 | 0.10633 | 0.64188 | 0.49871 | 2016-09-24 10:04 | 1.5706 | 1.5759 | 2016-09-24 10:05 | 1.5642 | 1.5694 | 2016-09-24 10:06 | 1.5579 | 1.563 | 2016-09-24 10:04 | 1.5674 | 1.5727 | 2016-09-24 10:06 | 1.5611 | 1.5662 | 0.63799 | 0.12367 | 0.64058 | 0.39238
51 | 29.7 | 99.3 | 24 | 10 | 8 | 0 | 9 | 16 | C | 83 | IZO | 48.3278 | 28.6303 | 89.8247 | 277.7927 | 295.8777 | 279.0718 | 263.8073 | 274.818 | 0.64702 | 0.2296 | 0.64957 | 0.1067 | 0.65211 | 0.49886 | 2016-09-24 10:08 | 1.5455 | 1.5504 | 2016-09-24 10:09 | 1.5395 | 1.5443 | 2016-09-24 10:10 | 1.5335 | 1.5382 | 2016-09-24 10:08 | 1.5425 | 1.5473 | 2016-09-24 10:10 | 1.5365 | 1.5412 | 0.6483 | 0.1229 | 0.65084 | 0.39217
50.4 | 29.3 | 97.9 | 24 | 10 | 12 | 0 | 9 | 16 | C | 83 | IZO | 47.7599 | 28.2247 | 88.5077 | 279.2261 | 296.5124 | 279.249 | 265.8583 | 274.8797 | 0.65716 | 0.23078 | 0.65966 | 0.10693 | 0.66216 | 0.49918 | 2016-09-24 10:12 | 1.5217 | 1.5262 | 2016-09-24 10:13 | 1.5159 | 1.5203 | 2016-09-24 10:14 | 1.5102 | 1.5145 | 2016-09-24 10:12 | 1.5188 | 1.5233 | 2016-09-24 10:14 | 1.5131 | 1.5174 | 0.65841 | 0.12386 | 0.66091 | 0.39225
49.6 | 28.8 | 96.6 | 24 | 10 | 16 | 0 | 9 | 16 | C | 83 | IZO | 47.0012 | 27.7169 | 87.2851 | 278.8675 | 294.8678 | 279.5357 | 266.4946 | 275.6547 | 0.6671 | 0.23048 | 0.66956 | 0.10633 | 0.672 | 0.49969 | 2016-09-24 10:16 | 1.499 | 1.5032 | 2016-09-24 10:17 | 1.4935 | 1.4976 | 2016-09-24 10:18 | 1.4881 | 1.4921 | 2016-09-24 10:16 | 1.4963 | 1.5004 | 2016-09-24 10:18 | 1.4908 | 1.4948 | 0.66833 | 0.12415 | 0.67078 | 0.39336
48.9 | 28.4 | 95.1 | 24 | 10 | 20 | 0 | 9 | 16 | C | 83 | IZO | 46.3358 | 27.3099 | 85.872 | 278.9776 | 294.6917 | 278.8647 | 266.8257 | 274.8588 | 0.67685 | 0.23058 | 0.67925 | 0.10627 | 0.68164 | 0.49849 | 2016-09-24 10:20 | 1.4774 | 1.4813 | 2016-09-24 10:21 | 1.4722 | 1.476 | 2016-09-24 10:22 | 1.467 | 1.4707 | 2016-09-24 10:20 | 1.4748 | 1.4786 | 2016-09-24 10:22 | 1.4696 | 1.4733 | 0.67805 | 0.12431 | 0.68045 | 0.39222
48.2 | 28 | 94 | 24 | 10 | 24 | 0 | 9 | 16 | C | 83 | IZO | 45.6687 | 26.9024 | 84.8356 | 278.8039 | 294.1312 | 279.4004 | 266.9509 | 275.6716 | 0.68639 | 0.23043 | 0.68874 | 0.10607 | 0.69108 | 0.49945 | 2016-09-24 10:24 | 1.4569 | 1.4604 | 2016-09-24 10:25 | 1.4519 | 1.4554 | 2016-09-24 10:26 | 1.447 | 1.4504 | 2016-09-24 10:24 | 1.4544 | 1.4579 | 2016-09-24 10:25 | 1.4495 | 1.4529 | 0.68757 | 0.12437 | 0.68992 | 0.39338
47.5 | 27.5 | 92.6 | 24 | 10 | 28 | 0 | 9 | 16 | C | 83 | IZO | 44.9998 | 26.3921 | 83.5186 | 278.3475 | 291.2136 | 278.6295 | 268.3992 | 275.4433 | 0.69573 | 0.23005 | 0.69803 | 0.10501 | 0.70032 | 0.49807 | 2016-09-24 10:28 | 1.4373 | 1.4406 | 2016-09-24 10:29 | 1.4326 | 1.4358 | 2016-09-24 10:30 | 1.4279 | 1.4311 | 2016-09-24 10:28 | 1.435 | 1.4382 | 2016-09-24 10:30 | 1.4303 | 1.4334 | 0.69688 | 0.12504 | 0.69917 | 0.39306
46.9 | 27.5 | 91.7 | 24 | 10 | 32 | 0 | 9 | 16 | C | 83 | IZO | 44.4248 | 26.3921 | 82.6742 | 278.428 | 297.8867 | 279.5009 | 263.3764 | 274.8488 | 0.70485 | 0.23012 | 0.7071 | 0.10742 | 0.70934 | 0.49963 | 2016-09-24 10:32 | 1.4187 | 1.4218 | 2016-09-24 10:33 | 1.4142 | 1.4172 | 2016-09-24 10:34 | 1.4098 | 1.4127 | 2016-09-24 10:32 | 1.4165 | 1.4195 | 2016-09-24 10:34 | 1.412 | 1.4149 | 0.70598 | 0.1227 | 0.70822 | 0.39221
46.4 | 27.1 | 90.6 | 24 | 10 | 36 | 0 | 9 | 16 | C | 83 | IZO | 43.9444 | 25.9832 | 81.6455 | 279.0851 | 296.2845 | 279.4488 | 265.7819 | 275.1885 | 0.71377 | 0.23066 | 0.71597 | 0.10684 | 0.71815 | 0.49954 | 2016-09-24 10:36 | 1.401 | 1.4038 | 2016-09-24 10:37 | 1.3967 | 1.3995 | 2016-09-24 10:38 | 1.3925 | 1.3952 | 2016-09-24 10:36 | 1.3989 | 1.4016 | 2016-09-24 10:38 | 1.3946 | 1.3973 | 0.71487 | 0.12382 | 0.71706 | 0.39269
45.8 | 26.8 | 89.6 | 24 | 10 | 40 | 0 | 9 | 16 | C | 83 | IZO | 43.3664 | 25.6762 | 80.7129 | 278.6725 | 296.3733 | 279.5884 | 264.9809 | 275.341 | 0.72247 | 0.23032 | 0.72461 | 0.10688 | 0.72674 | 0.49979 | 2016-09-24 10:40 | 1.3841 | 1.3867 | 2016-09-24 10:41 | 1.38 | 1.3826 | 2016-09-24 10:42 | 1.376 | 1.3785 | 2016-09-24 10:40 | 1.3821 | 1.3846 | 2016-09-24 10:41 | 1.378 | 1.3805 | 0.72355 | 0.12345 | 0.72568 | 0.39291
45.4 | 26.6 | 88.6 | 24 | 10 | 44 | 0 | 9 | 16 | C | 83 | IZO | 42.9805 | 25.4713 | 79.7824 | 279.7229 | 298.2367 | 279.5485 | 265.4016 | 274.8203 | 0.73096 | 0.23119 | 0.73304 | 0.10755 | 0.73512 | 0.49971 | 2016-09-24 10:44 | 1.3681 | 1.3705 | 2016-09-24 10:45 | 1.3642 | 1.3665 | 2016-09-24 10:46 | 1.3603 | 1.3626 | 2016-09-24 10:44 | 1.3661 | 1.3685 | 2016-09-24 10:46 | 1.3622 | 1.3646 | 0.732 | 0.12364 | 0.73408 | 0.39217
44.8 | 26.2 | 87.8 | 24 | 10 | 48 | 0 | 9 | 16 | C | 83 | IZO | 42.4005 | 25.0612 | 79.0391 | 278.8449 | 295.6285 | 280.1021 | 265.8626 | 276.173 | 0.73922 | 0.23047 | 0.74125 | 0.10661 | 0.74327 | 0.5007 | 2016-09-24 10:48 | 1.3528 | 1.355 | 2016-09-24 10:49 | 1.3491 | 1.3512 | 2016-09-24 10:50 | 1.3454 | 1.3475 | 2016-09-24 10:48 | 1.3509 | 1.3531 | 2016-09-24 10:50 | 1.3472 | 1.3494 | 0.74024 | 0.12386 | 0.74226 | 0.3941
44.5 | 25.9 | 86.8 | 24 | 10 | 52 | 0 | 9 | 16 | C | 83 | IZO | 42.1101 | 24.7532 | 78.1112 | 280.3549 | 294.8023 | 279.7135 | 269.1806 | 275.895 | 0.74726 | 0.23171 | 0.74924 | 0.10631 | 0.7512 | 0.50001 | 2016-09-24 10:52 | 1.3382 | 1.3403 | 2016-09-24 10:53 | 1.3347 | 1.3367 | 2016-09-24 10:54 | 1.3312 | 1.3332 | 2016-09-24 10:52 | 1.3364 | 1.3385 | 2016-09-24 10:54 | 1.3329 | 1.3349 | 0.74825 | 0.12541 | 0.75022 | 0.3937
43.9 | 25.7 | 85.9 | 24 | 10 | 56 | 0 | 9 | 16 | C | 83 | IZO | 41.5283 | 24.5477 | 77.2769 | 279.0313 | 295.8366 | 279.5418 | 266.0317 | 275.4186 | 0.75508 | 0.23062 | 0.757 | 0.10668 | 0.7589 | 0.4997 | 2016-09-24 10:56 | 1.3244 | 1.3263 | 2016-09-24 10:57 | 1.321 | 1.3229 | 2016-09-24 10:58 | 1.3177 | 1.3195 | 2016-09-24 10:56 | 1.3227 | 1.3246 | 2016-09-24 10:57 | 1.3193 | 1.3212 | 0.75604 | 0.12394 | 0.75795 | 0.39302
43.5 | 25.5 | 85.1 | 24 | 11 | 0 | 0 | 9 | 16 | C | 83 | IZO | 41.1396 | 24.3421 | 76.5357 | 279.2669 | 296.624 | 279.5991 | 265.84 | 275.2916 | 0.76267 | 0.23081 | 0.76453 | 0.10697 | 0.76637 | 0.49981 | 2016-09-24 11:00 | 1.3112 | 1.3129 | 2016-09-24 11:01 | 1.308 | 1.3097 | 2016-09-24 11:02 | 1.3048 | 1.3065 | 2016-09-24 11:01 | 1.3096 | 1.3113 | 2016-09-24 11:02 | 1.3064 | 1.3081 | 0.7636 | 0.12385 | 0.76545 | 0.39284
43.1 | 25.2 | 84.2 | 24 | 11 | 4 | 0 | 9 | 16 | C | 83 | IZO | 40.7503 | 24.0335 | 75.7025 | 279.3111 | 294.9641 | 279.0978 | 267.203 | 275.0831 | 0.77002 | 0.23085 | 0.77183 | 0.10637 | 0.77361 | 0.49891 | 2016-09-24 11:04 | 1.2987 | 1.3003 | 2016-09-24 11:05 | 1.2956 | 1.2972 | 2016-09-24 11:06 | 1.2926 | 1.2942 | 2016-09-24 11:04 | 1.2971 | 1.2987 | 2016-09-24 11:06 | 1.2941 | 1.2957 | 0.77093 | 0.12448 | 0.77272 | 0.39254
42.7 | 24.9 | 83.4 | 24 | 11 | 8 | 0 | 9 | 16 | C | 83 | IZO | 40.3604 | 23.7245 | 74.9633 | 279.1658 | 293.0184 | 278.8415 | 268.451 | 275.2539 | 0.77715 | 0.23073 | 0.77889 | 0.10567 | 0.78062 | 0.49845 | 2016-09-24 11:08 | 1.2868 | 1.2882 | 2016-09-24 11:09 | 1.2839 | 1.2853 | 2016-09-24 11:10 | 1.281 | 1.2825 | 2016-09-24 11:08 | 1.2853 | 1.2868 | 2016-09-24 11:10 | 1.2824 | 1.2839 | 0.77802 | 0.12507 | 0.77976 | 0.39279
42.4 | 25.1 | 83.1 | 24 | 11 | 12 | 0 | 9 | 16 | C | 83 | IZO | 40.0675 | 23.9305 | 74.6864 | 279.7597 | 302.0177 | 280.4652 | 262.5387 | 275.0138 | 0.78404 | 0.23122 | 0.78572 | 0.10891 | 0.78739 | 0.50135 | 2016-09-24 11:12 | 1.2754 | 1.2768 | 2016-09-24 11:13 | 1.2727 | 1.2741 | 2016-09-24 11:14 | 1.27 | 1.2713 | 2016-09-24 11:12 | 1.2741 | 1.2754 | 2016-09-24 11:13 | 1.2714 | 1.2727 | 0.78488 | 0.12231 | 0.78656 | 0.39244
42 | 24.9 | 82.5 | 24 | 11 | 16 | 0 | 9 | 16 | C | 83 | IZO | 39.6764 | 23.7245 | 74.133 | 279.2505 | 301.8826 | 280.7403 | 261.7396 | 275.3926 | 0.79069 | 0.2308 | 0.79232 | 0.10886 | 0.79393 | 0.50185 | 2016-09-24 11:16 | 1.2647 | 1.266 | 2016-09-24 11:17 | 1.2621 | 1.2634 | 2016-09-24 11:18 | 1.2596 | 1.2608 | 2016-09-24 11:17 | 1.2634 | 1.2647 | 2016-09-24 11:18 | 1.2608 | 1.2621 | 0.7915 | 0.12194 | 0.79312 | 0.39298
41.7 | 24.7 | 81.8 | 24 | 11 | 20 | 0 | 9 | 16 | C | 83 | IZO | 39.3826 | 23.5184 | 73.4877 | 279.5027 | 301.512 | 280.4647 | 262.4737 | 275.141 | 0.7971 | 0.23101 | 0.79867 | 0.10873 | 0.80022 | 0.50135 | 2016-09-24 11:20 | 1.2545 | 1.2557 | 2016-09-24 11:21 | 1.2521 | 1.2532 | 2016-09-24 11:22 | 1.2497 | 1.2508 | 2016-09-24 11:20 | 1.2533 | 1.2545 | 2016-09-24 11:22 | 1.2509 | 1.252 | 0.79789 | 0.12228 | 0.79945 | 0.39262

## IdD083_star
This Page uses the lamp tests and the langley plot results in the limited mu range from sheet IdD083 | [21 empty] | the correction values from ID083 are applied to the NC
Dobson 083 at IZO 2016.Sep.24 Langley Plot Results | [15 empty] | 109
SLC083Na | -0.9 | SLC083Nc | -0.7 | [9 empty] | SLC083Nd | -0.7
D083 LP adj. A:-0.01 C:-0.11 D:-0.13 Nad:0.11 | [20 empty] | LPC_C_83 |  | LPC_D_83 |  | LPC_A_83
[15 empty] | 275.4705 | 284.433 | 277.1703 | 268.5578 | 275.3244 | All Obs | -0.10657 |  | -0.12532 |  | -0.011183 | 0.11414 | 0.018752
D083RC | D083RD | D083RA | day | h | m | s | m | year |  |  |  | D083NC | D083ND | D083NA | D083_TO3_C* | D083_TO3_D* | D083_TO3_A* | D083_TO3_CD* | D083_TO3_AD* | 1/mu.C | P*C_83 | 1/Mu.D | P*D_83 | 1/Mu.A | P*A_83 | lst.C | Mu.C | AirM.C | lst.D | Mu.D | AirM.D | lst.A | Mu.A | AirM.A | lst.CD | Mu.CD | AirM.CD | lst.AD | Mu.AD | AirM.AD
97.9 | 55.4 | 181 | 24 | 8 | 28 | 0 | 9 | 16 | C | 83 | IZO | 91.3643 | 53.0739 | 167.4449 | 275.2653 | 284.1453 | 276.904 | 268.48 | 275.043 | 0.34153 | 0.22751 | 0.34494 | 0.10247 | 0.34836 | 0.49499 | 2016-09-24 08:28 | 2.928 | 2.9879 | 2016-09-24 08:29 | 2.899 | 2.9569 | 2016-09-24 08:30 | 2.8706 | 2.9265 | 2016-09-24 08:28 | 2.9135 | 2.9723 | 2016-09-24 08:29 | 2.8847 | 2.9416
93.9 | 53.1 | 174.5 | 24 | 8 | 32 | 0 | 9 | 16 | C | 83 | IZO | 87.6565 | 50.9228 | 161.0817 | 274.6037 | 283.2094 | 276.8554 | 268.0206 | 275.2218 | 0.35517 | 0.22696 | 0.35858 | 0.10213 | 0.36197 | 0.4949 | 2016-09-24 08:32 | 2.8155 | 2.8679 | 2016-09-24 08:33 | 2.7888 | 2.8395 | 2016-09-24 08:34 | 2.7626 | 2.8117 | 2016-09-24 08:33 | 2.8021 | 2.8536 | 2016-09-24 08:34 | 2.7757 | 2.8255
90.7 | 51.1 | 168.5 | 24 | 8 | 36 | 0 | 9 | 16 | C | 83 | IZO | 84.6913 | 49.0376 | 155.2171 | 275.9372 | 283.2471 | 276.8097 | 270.3485 | 275.1578 | 0.36875 | 0.22806 | 0.37213 | 0.10214 | 0.37551 | 0.49482 | 2016-09-24 08:36 | 2.7119 | 2.7578 | 2016-09-24 08:37 | 2.6872 | 2.7318 | 2016-09-24 08:38 | 2.663 | 2.7062 | 2016-09-24 08:36 | 2.6995 | 2.7447 | 2016-09-24 08:38 | 2.6751 | 2.7189
87.2 | 49.2 | 163 | 24 | 8 | 40 | 0 | 9 | 16 | C | 83 | IZO | 81.4448 | 47.2358 | 149.8341 | 274.8984 | 282.613 | 276.8458 | 268.9892 | 275.3656 | 0.38225 | 0.2272 | 0.38561 | 0.10191 | 0.38897 | 0.49488 | 2016-09-24 08:40 | 2.6161 | 2.6567 | 2016-09-24 08:41 | 2.5933 | 2.6327 | 2016-09-24 08:42 | 2.5709 | 2.6091 | 2016-09-24 08:40 | 2.6046 | 2.6446 | 2016-09-24 08:42 | 2.5821 | 2.6208
84.2 | 47.6 | 158 | 24 | 8 | 44 | 0 | 9 | 16 | C | 83 | IZO | 78.6711 | 45.7097 | 144.9709 | 274.9639 | 283.5375 | 277.1611 | 268.3836 | 275.529 | 0.39566 | 0.22726 | 0.399 | 0.10225 | 0.40233 | 0.49545 | 2016-09-24 08:44 | 2.5274 | 2.5634 | 2016-09-24 08:45 | 2.5063 | 2.5412 | 2016-09-24 08:46 | 2.4855 | 2.5194 | 2016-09-24 08:44 | 2.5168 | 2.5522 | 2016-09-24 08:46 | 2.4959 | 2.5303
81.5 | 46.1 | 153.3 | 24 | 8 | 48 | 0 | 9 | 16 | C | 83 | IZO | 76.186 | 44.2712 | 140.4208 | 275.4576 | 284.1842 | 277.3944 | 268.7534 | 275.6595 | 0.40898 | 0.22767 | 0.41229 | 0.10248 | 0.4156 | 0.49586 | 2016-09-24 08:48 | 2.4451 | 2.4771 | 2016-09-24 08:49 | 2.4255 | 2.4566 | 2016-09-24 08:50 | 2.4062 | 2.4364 | 2016-09-24 08:49 | 2.4353 | 2.4668 | 2016-09-24 08:50 | 2.4158 | 2.4464
78.8 | 44.6 | 148.9 | 24 | 8 | 52 | 0 | 9 | 16 | C | 83 | IZO | 73.7027 | 42.8255 | 136.1572 | 275.0571 | 283.539 | 277.555 | 268.5378 | 276.0253 | 0.42219 | 0.22734 | 0.42548 | 0.10225 | 0.42876 | 0.49615 | 2016-09-24 08:52 | 2.3686 | 2.3971 | 2016-09-24 08:53 | 2.3503 | 2.3781 | 2016-09-24 08:54 | 2.3323 | 2.3593 | 2016-09-24 08:52 | 2.3594 | 2.3876 | 2016-09-24 08:54 | 2.3413 | 2.3686
76.3 | 43.3 | 144.5 | 24 | 8 | 56 | 0 | 9 | 16 | C | 83 | IZO | 71.3996 | 41.5668 | 131.9399 | 274.7097 | 283.9634 | 277.1161 | 267.5888 | 275.3695 | 0.4353 | 0.22705 | 0.43856 | 0.1024 | 0.44181 | 0.49537 | 2016-09-24 08:56 | 2.2973 | 2.3228 | 2016-09-24 08:57 | 2.2802 | 2.3051 | 2016-09-24 08:58 | 2.2634 | 2.2876 | 2016-09-24 08:56 | 2.2887 | 2.3139 | 2016-09-24 08:58 | 2.2718 | 2.2963
74.2 | 42 | 140.6 | 24 | 9 | 0 | 0 | 9 | 16 | C | 83 | IZO | 69.4643 | 40.3024 | 128.2173 | 275.5227 | 283.2559 | 277.2328 | 269.5752 | 275.6956 | 0.44829 | 0.22772 | 0.45152 | 0.10215 | 0.45475 | 0.49558 | 2016-09-24 09:00 | 2.2307 | 2.2536 | 2016-09-24 09:01 | 2.2147 | 2.2371 | 2016-09-24 09:02 | 2.199 | 2.2208 | 2016-09-24 09:00 | 2.2227 | 2.2453 | 2016-09-24 09:02 | 2.2068 | 2.2289
72 | 41 | 137.1 | 24 | 9 | 4 | 0 | 9 | 16 | C | 83 | IZO | 67.4369 | 39.3257 | 124.8824 | 275.1082 | 285.1969 | 277.7363 | 267.3343 | 275.8369 | 0.46117 | 0.22738 | 0.46437 | 0.10285 | 0.46756 | 0.49648 | 2016-09-24 09:04 | 2.1684 | 2.189 | 2016-09-24 09:05 | 2.1535 | 2.1736 | 2016-09-24 09:06 | 2.1388 | 2.1584 | 2016-09-24 09:05 | 2.1609 | 2.1813 | 2016-09-24 09:06 | 2.1461 | 2.1659
70 | 39.8 | 133.5 | 24 | 9 | 8 | 0 | 9 | 16 | C | 83 | IZO | 65.5953 | 38.1486 | 121.4448 | 275.0215 | 283.6434 | 277.3962 | 268.3802 | 275.8044 | 0.47392 | 0.22731 | 0.47709 | 0.10228 | 0.48025 | 0.49587 | 2016-09-24 09:08 | 2.1101 | 2.1287 | 2016-09-24 09:09 | 2.096 | 2.1142 | 2016-09-24 09:10 | 2.0823 | 2.1 | 2016-09-24 09:08 | 2.103 | 2.1214 | 2016-09-24 09:10 | 2.0891 | 2.1071
68.1 | 38.9 | 130.2 | 24 | 9 | 12 | 0 | 9 | 16 | C | 83 | IZO | 63.8463 | 37.262 | 118.291 | 274.8055 | 285.056 | 277.2602 | 266.9011 | 275.2778 | 0.48654 | 0.22713 | 0.48968 | 0.10279 | 0.4928 | 0.49562 | 2016-09-24 09:12 | 2.0553 | 2.0722 | 2016-09-24 09:13 | 2.0422 | 2.0586 | 2016-09-24 09:14 | 2.0292 | 2.0453 | 2016-09-24 09:12 | 2.0487 | 2.0654 | 2016-09-24 09:14 | 2.0357 | 2.0519
66.5 | 38.1 | 127.3 | 24 | 9 | 16 | 0 | 9 | 16 | C | 83 | IZO | 62.3718 | 36.4712 | 115.5258 | 275.6063 | 287.0347 | 277.6884 | 266.788 | 275.3149 | 0.49903 | 0.22779 | 0.50213 | 0.10351 | 0.50522 | 0.49639 | 2016-09-24 09:16 | 2.0039 | 2.0192 | 2016-09-24 09:17 | 1.9915 | 2.0064 | 2016-09-24 09:18 | 1.9793 | 1.9939 | 2016-09-24 09:16 | 1.9977 | 2.0128 | 2016-09-24 09:18 | 1.9854 | 2.0001
55.9 | 32.2 | 107.9 | 24 | 9 | 48 | 0 | 9 | 16 | C | 83 | IZO | 52.5352 | 30.5791 | 97.2786 | 276.5887 | 285.5289 | 277.3331 | 269.6853 | 275.2539 | 0.5935 | 0.2286 | 0.59629 | 0.10296 | 0.59906 | 0.49575 | 2016-09-24 09:48 | 1.6849 | 1.6923 | 2016-09-24 09:49 | 1.677 | 1.6843 | 2016-09-24 09:50 | 1.6693 | 1.6764 | 2016-09-24 09:48 | 1.681 | 1.6883 | 2016-09-24 09:50 | 1.6732 | 1.6803
55 | 31.5 | 106.1 | 24 | 9 | 52 | 0 | 9 | 16 | C | 83 | IZO | 51.6935 | 29.8743 | 95.605 | 277.4952 | 283.0674 | 277.6096 | 273.1979 | 276.2227 | 0.60457 | 0.22935 | 0.60731 | 0.10208 | 0.61004 | 0.49625 | 2016-09-24 09:52 | 1.6541 | 1.6609 | 2016-09-24 09:53 | 1.6466 | 1.6533 | 2016-09-24 09:54 | 1.6392 | 1.6458 | 2016-09-24 09:52 | 1.6503 | 1.657 | 2016-09-24 09:54 | 1.6429 | 1.6495
53.7 | 31 | 104.2 | 24 | 9 | 56 | 0 | 9 | 16 | C | 83 | IZO | 50.4744 | 29.37 | 93.8339 | 275.2538 | 283.4634 | 277.2462 | 268.9143 | 275.6675 | 0.61546 | 0.2275 | 0.61815 | 0.10222 | 0.62083 | 0.4956 | 2016-09-24 09:56 | 1.6248 | 1.6311 | 2016-09-24 09:57 | 1.6177 | 1.6238 | 2016-09-24 09:58 | 1.6107 | 1.6167 | 2016-09-24 09:56 | 1.6213 | 1.6274 | 2016-09-24 09:58 | 1.6142 | 1.6203
52.8 | 30.5 | 102.4 | 24 | 10 | 0 | 0 | 9 | 16 | C | 83 | IZO | 49.6275 | 28.865 | 92.1476 | 275.4012 | 283.393 | 276.873 | 269.2294 | 275.2181 | 0.62617 | 0.22762 | 0.62881 | 0.10219 | 0.63145 | 0.49493 | 2016-09-24 10:00 | 1.597 | 1.6028 | 2016-09-24 10:01 | 1.5903 | 1.5959 | 2016-09-24 10:02 | 1.5837 | 1.5892 | 2016-09-24 10:00 | 1.5936 | 1.5993 | 2016-09-24 10:02 | 1.587 | 1.5925
51.9 | 30.1 | 100.8 | 24 | 10 | 4 | 0 | 9 | 16 | C | 83 | IZO | 48.7784 | 28.4605 | 90.6445 | 275.2006 | 284.6546 | 276.8616 | 267.8964 | 274.8853 | 0.63669 | 0.22745 | 0.63929 | 0.10265 | 0.64188 | 0.49491 | 2016-09-24 10:04 | 1.5706 | 1.5759 | 2016-09-24 10:05 | 1.5642 | 1.5694 | 2016-09-24 10:06 | 1.5579 | 1.563 | 2016-09-24 10:04 | 1.5674 | 1.5727 | 2016-09-24 10:06 | 1.5611 | 1.5662
51 | 29.7 | 99.3 | 24 | 10 | 8 | 0 | 9 | 16 | C | 83 | IZO | 47.9279 | 28.0554 | 89.2336 | 274.662 | 285.5204 | 276.9152 | 266.2701 | 274.7338 | 0.64702 | 0.22701 | 0.64957 | 0.10296 | 0.65211 | 0.49501 | 2016-09-24 10:08 | 1.5455 | 1.5504 | 2016-09-24 10:09 | 1.5395 | 1.5443 | 2016-09-24 10:10 | 1.5335 | 1.5382 | 2016-09-24 10:08 | 1.5425 | 1.5473 | 2016-09-24 10:10 | 1.5365 | 1.5412
50.4 | 29.3 | 97.9 | 24 | 10 | 12 | 0 | 9 | 16 | C | 83 | IZO | 47.3599 | 27.6497 | 87.9166 | 276.0463 | 285.9942 | 277.0592 | 268.3586 | 274.7946 | 0.65716 | 0.22815 | 0.65966 | 0.10313 | 0.66216 | 0.49527 | 2016-09-24 10:12 | 1.5217 | 1.5262 | 2016-09-24 10:13 | 1.5159 | 1.5203 | 2016-09-24 10:14 | 1.5102 | 1.5145 | 2016-09-24 10:12 | 1.5188 | 1.5233 | 2016-09-24 10:14 | 1.5131 | 1.5174
49.6 | 28.8 | 96.6 | 24 | 10 | 16 | 0 | 9 | 16 | C | 83 | IZO | 46.6013 | 27.1419 | 86.6939 | 275.6396 | 284.1919 | 277.3134 | 269.0316 | 275.5687 | 0.6671 | 0.22782 | 0.66956 | 0.10248 | 0.672 | 0.49572 | 2016-09-24 10:16 | 1.499 | 1.5032 | 2016-09-24 10:17 | 1.4935 | 1.4976 | 2016-09-24 10:18 | 1.4881 | 1.4921 | 2016-09-24 10:16 | 1.4963 | 1.5004 | 2016-09-24 10:18 | 1.4908 | 1.4948
48.9 | 28.4 | 95.1 | 24 | 10 | 20 | 0 | 9 | 16 | C | 83 | IZO | 45.9359 | 26.7349 | 85.2808 | 275.7025 | 283.8612 | 276.6105 | 269.3987 | 274.7719 | 0.67685 | 0.22787 | 0.67925 | 0.10236 | 0.68164 | 0.49446 | 2016-09-24 10:20 | 1.4774 | 1.4813 | 2016-09-24 10:21 | 1.4722 | 1.476 | 2016-09-24 10:22 | 1.467 | 1.4707 | 2016-09-24 10:20 | 1.4748 | 1.4786 | 2016-09-24 10:22 | 1.4696 | 1.4733
48.2 | 28 | 94 | 24 | 10 | 24 | 0 | 9 | 16 | C | 83 | IZO | 45.2688 | 26.3274 | 84.2445 | 275.4826 | 283.1494 | 277.1149 | 269.5591 | 275.5838 | 0.68639 | 0.22769 | 0.68874 | 0.10211 | 0.69108 | 0.49536 | 2016-09-24 10:24 | 1.4569 | 1.4604 | 2016-09-24 10:25 | 1.4519 | 1.4554 | 2016-09-24 10:26 | 1.447 | 1.4504 | 2016-09-24 10:24 | 1.4544 | 1.4579 | 2016-09-24 10:25 | 1.4495 | 1.4529
47.5 | 27.5 | 92.6 | 24 | 10 | 28 | 0 | 9 | 16 | C | 83 | IZO | 44.5999 | 25.8171 | 82.9274 | 274.9811 | 280.0837 | 276.3135 | 271.0419 | 275.3547 | 0.69573 | 0.22727 | 0.69803 | 0.101 | 0.70032 | 0.49393 | 2016-09-24 10:28 | 1.4373 | 1.4406 | 2016-09-24 10:29 | 1.4326 | 1.4358 | 2016-09-24 10:30 | 1.4279 | 1.4311 | 2016-09-24 10:28 | 1.435 | 1.4382 | 2016-09-24 10:30 | 1.4303 | 1.4334
46.9 | 27.5 | 91.7 | 24 | 10 | 32 | 0 | 9 | 16 | C | 83 | IZO | 44.0249 | 25.8171 | 82.083 | 275.0174 | 286.6121 | 277.1551 | 266.0528 | 274.7594 | 0.70485 | 0.2273 | 0.7071 | 0.10336 | 0.70934 | 0.49544 | 2016-09-24 10:32 | 1.4187 | 1.4218 | 2016-09-24 10:33 | 1.4142 | 1.4172 | 2016-09-24 10:34 | 1.4098 | 1.4127 | 2016-09-24 10:32 | 1.4165 | 1.4195 | 2016-09-24 10:34 | 1.412 | 1.4149
46.4 | 27.1 | 90.6 | 24 | 10 | 36 | 0 | 9 | 16 | C | 83 | IZO | 43.5444 | 25.4082 | 81.0543 | 275.6313 | 284.8686 | 277.0739 | 268.4911 | 275.0983 | 0.71377 | 0.22781 | 0.71597 | 0.10273 | 0.71815 | 0.49529 | 2016-09-24 10:36 | 1.401 | 1.4038 | 2016-09-24 10:37 | 1.3967 | 1.3995 | 2016-09-24 10:38 | 1.3925 | 1.3952 | 2016-09-24 10:36 | 1.3989 | 1.4016 | 2016-09-24 10:38 | 1.3946 | 1.3973
45.8 | 26.8 | 89.6 | 24 | 10 | 40 | 0 | 9 | 16 | C | 83 | IZO | 42.9665 | 25.1012 | 80.1218 | 275.1767 | 284.8195 | 277.185 | 267.7221 | 275.2501 | 0.72247 | 0.22743 | 0.72461 | 0.10271 | 0.72674 | 0.49549 | 2016-09-24 10:40 | 1.3841 | 1.3867 | 2016-09-24 10:41 | 1.38 | 1.3826 | 2016-09-24 10:42 | 1.376 | 1.3785 | 2016-09-24 10:40 | 1.3821 | 1.3846 | 2016-09-24 10:41 | 1.378 | 1.3805
45.4 | 26.6 | 88.6 | 24 | 10 | 44 | 0 | 9 | 16 | C | 83 | IZO | 42.5806 | 24.8963 | 79.1912 | 276.186 | 286.5485 | 277.1174 | 268.174 | 274.7286 | 0.73096 | 0.22827 | 0.73304 | 0.10333 | 0.73512 | 0.49537 | 2016-09-24 10:44 | 1.3681 | 1.3705 | 2016-09-24 10:45 | 1.3642 | 1.3665 | 2016-09-24 10:46 | 1.3603 | 1.3626 | 2016-09-24 10:44 | 1.3661 | 1.3685 | 2016-09-24 10:46 | 1.3622 | 1.3646
44.8 | 26.2 | 87.8 | 24 | 10 | 48 | 0 | 9 | 16 | C | 83 | IZO | 42.0006 | 24.4862 | 78.4479 | 275.268 | 283.8094 | 277.6441 | 268.6654 | 276.0806 | 0.73922 | 0.22751 | 0.74125 | 0.10234 | 0.74327 | 0.49631 | 2016-09-24 10:48 | 1.3528 | 1.355 | 2016-09-24 10:49 | 1.3491 | 1.3512 | 2016-09-24 10:50 | 1.3454 | 1.3475 | 2016-09-24 10:48 | 1.3509 | 1.3531 | 2016-09-24 10:50 | 1.3472 | 1.3494
44.5 | 25.9 | 86.8 | 24 | 10 | 52 | 0 | 9 | 16 | C | 83 | IZO | 41.7102 | 24.1782 | 77.52 | 276.7391 | 282.8559 | 277.2292 | 272.0129 | 275.8019 | 0.74726 | 0.22873 | 0.74924 | 0.102 | 0.7512 | 0.49557 | 2016-09-24 10:52 | 1.3382 | 1.3403 | 2016-09-24 10:53 | 1.3347 | 1.3367 | 2016-09-24 10:54 | 1.3312 | 1.3332 | 2016-09-24 10:52 | 1.3364 | 1.3385 | 2016-09-24 10:54 | 1.3329 | 1.3349
43.9 | 25.7 | 85.9 | 24 | 10 | 56 | 0 | 9 | 16 | C | 83 | IZO | 41.1283 | 23.9727 | 76.6857 | 275.3777 | 283.7665 | 277.032 | 268.8927 | 275.3249 | 0.75508 | 0.2276 | 0.757 | 0.10233 | 0.7589 | 0.49522 | 2016-09-24 10:56 | 1.3244 | 1.3263 | 2016-09-24 10:57 | 1.321 | 1.3229 | 2016-09-24 10:58 | 1.3177 | 1.3195 | 2016-09-24 10:56 | 1.3227 | 1.3246 | 2016-09-24 10:57 | 1.3193 | 1.3212
43.5 | 25.5 | 85.1 | 24 | 11 | 0 | 0 | 9 | 16 | C | 83 | IZO | 40.7397 | 23.7671 | 75.9445 | 275.5766 | 284.4338 | 277.0647 | 268.7287 | 275.1972 | 0.76267 | 0.22776 | 0.76453 | 0.10257 | 0.76637 | 0.49527 | 2016-09-24 11:00 | 1.3112 | 1.3129 | 2016-09-24 11:01 | 1.308 | 1.3097 | 2016-09-24 11:02 | 1.3048 | 1.3065 | 2016-09-24 11:01 | 1.3096 | 1.3113 | 2016-09-24 11:02 | 1.3064 | 1.3081
43.1 | 25.2 | 84.2 | 24 | 11 | 4 | 0 | 9 | 16 | C | 83 | IZO | 40.3504 | 23.4585 | 75.1113 | 275.5852 | 282.6576 | 276.5394 | 270.1187 | 274.9882 | 0.77002 | 0.22777 | 0.77183 | 0.10193 | 0.77361 | 0.49434 | 2016-09-24 11:04 | 1.2987 | 1.3003 | 2016-09-24 11:05 | 1.2956 | 1.2972 | 2016-09-24 11:06 | 1.2926 | 1.2942 | 2016-09-24 11:04 | 1.2971 | 1.2987 | 2016-09-24 11:06 | 1.2941 | 1.2957
42.7 | 24.9 | 83.4 | 24 | 11 | 8 | 0 | 9 | 16 | C | 83 | IZO | 39.9604 | 23.1495 | 74.3721 | 275.4054 | 280.5991 | 276.26 | 271.3928 | 275.1584 | 0.77715 | 0.22762 | 0.77889 | 0.10119 | 0.78062 | 0.49384 | 2016-09-24 11:08 | 1.2868 | 1.2882 | 2016-09-24 11:09 | 1.2839 | 1.2853 | 2016-09-24 11:10 | 1.281 | 1.2825 | 2016-09-24 11:08 | 1.2853 | 1.2868 | 2016-09-24 11:10 | 1.2824 | 1.2839
42.4 | 25.1 | 83.1 | 24 | 11 | 12 | 0 | 9 | 16 | C | 83 | IZO | 39.6676 | 23.3555 | 74.0953 | 275.966 | 289.4896 | 277.8613 | 265.5056 | 274.9177 | 0.78404 | 0.22809 | 0.78572 | 0.10439 | 0.78739 | 0.4967 | 2016-09-24 11:12 | 1.2754 | 1.2768 | 2016-09-24 11:13 | 1.2727 | 1.2741 | 2016-09-24 11:14 | 1.27 | 1.2713 | 2016-09-24 11:12 | 1.2741 | 1.2754 | 2016-09-24 11:13 | 1.2714 | 1.2727
42 | 24.9 | 82.5 | 24 | 11 | 16 | 0 | 9 | 16 | C | 83 | IZO | 39.2764 | 23.1495 | 73.5418 | 275.4246 | 289.2493 | 278.1148 | 264.7308 | 275.296 | 0.79069 | 0.22764 | 0.79232 | 0.10431 | 0.79393 | 0.49715 | 2016-09-24 11:16 | 1.2647 | 1.266 | 2016-09-24 11:17 | 1.2621 | 1.2634 | 2016-09-24 11:18 | 1.2596 | 1.2608 | 2016-09-24 11:17 | 1.2634 | 1.2647 | 2016-09-24 11:18 | 1.2608 | 1.2621
41.7 | 24.7 | 81.8 | 24 | 11 | 20 | 0 | 9 | 16 | C | 83 | IZO | 38.9826 | 22.9434 | 72.8966 | 275.6457 | 288.7774 | 277.8183 | 265.4882 | 275.044 | 0.7971 | 0.22782 | 0.79867 | 0.10414 | 0.80022 | 0.49662 | 2016-09-24 11:20 | 1.2545 | 1.2557 | 2016-09-24 11:21 | 1.2521 | 1.2532 | 2016-09-24 11:22 | 1.2497 | 1.2508 | 2016-09-24 11:20 | 1.2533 | 1.2545 | 2016-09-24 11:22 | 1.2509 | 1.252

## param
-0.9 | 1e-05 |  | -1.9 | 0
0 | 1 |  | -1 | 1
0.9 | 1e-05 |  | 1 | 1
|  |  | 1.9 | 0
304.6 | 1e-07 | 310.55 | 1e-06 | 316.7 | 1e-06 | 323.5 | 1e-06 | 323.1 | 1e-06 | 330.5 | 1e-06 | 338 | 1e-06
305.5 | 1000 | 311.45 | 1000 | 317.6 | 1000 | 324.4 | 1000 | 324 | 1000 | 331.4 | 1000 | 338.9 | 1000
306.4 | 1e-07 | 312.35 | 1e-06 | 318.5 | 1e-06 | 326.4 | 1000 | 326 | 1000 | 333.4 | 1000 | 340.9 | 1000
[6 empty] | 327.3 | 1e-06 | 326.9 | 1e-06 | 334.3 | 1e-06 | 341.8 | 1e-06
1.1111
304.6 | 0.001 | 310.55 | 0.001 | 316.7 | 0.001 | 323.5 | 0 | 323.1 | 0.001 | 330.5 | 0.001 | 338 | 1e-07
305.5 | 100 | 311.45 | 100 | 317.6 | 100 | 324.4 | 100 | 324 | 100 | 331.4 | 100 | 338.9 | 100
306.4 | 0.001 | 312.35 | 0.001 | 318.5 | 0.001 | 326.4 | 100 | 326 | 100 | 333.4 | 100 | 340.9 | 100
[6 empty] | 327.3 | 0 | 326.9 | 0.001 | 334.3 | 0.001 | 341.8 | 1e-07
58 | 0
92 | 0
code | number | name | Dobson | earth_radius | LONG | Hemisphere | degs | min | sec | LAT | Hemisphere | deg | min | sec | TimeZone | stn Pressure | elevation | Ozone_Height | LongRadian | LatRadian
IZO | 999 | Izana Observatory,Spain | 83 | 6373593 | 16.4973 | W | 16 | 29 | 25 | 28.2986 | N | 28 | 17 | 55 | 0 | 770 | 2360 | 26000 | 0.28781 | 0.4939
=IF( C13<90,(earth_radius+Ozone_height)/(SQRT(((earth_radius+Ozone_height)^2)-((earth_radius+Elevation)^2)*(SIN(RADIANS(C13)))^2)),-90000)
Public Function AirMass#(CosZenAng As Double) | [14 empty] | p0 | 0.75993
AirMass# = 1# / (CosZenAng + 0.15 / (90 - RadsToDegs#(ArcCos#(CosZenAng)) + 3.885) ^ 1.253)
End Function
Ozone Absorption and Molecular Scattering coefficients for Use
with Dobson ozone Spectrophotometers Beginning 1 January 1992
Wavelength | ALFA |  | BETA | d(BETA)/d(ALFA) | [5 empty] | IUP Corr-factor | B&BP Corr_factor | D&M |  |  | α_Komhyr r adj
(nm) | (atm-cm)-1 |  | (atm)-1 | (atm-cm/atm) |  |  |  |  | A_ pair | 0.9898 | 0.9969 | 0.9889 |  |  | 1.806
[9 empty] | C_pair | 0.9922 | 0.9983 | 0.9845 |  |  | 0.833
305 | 0.489 | [7 empty] | D_pair | 0.9642 | 0.9825 | 0.9562 |  |  | 0.374
325 | 0.375 | [7 empty] | AD_pair | 0.9965 | 1.0006 | 0.9975 |  |  | 1.432
Alfa_A | 1.7876 | Beta_A | 114 | 63 | 1000 |  |  |  | CD_pair | 1.015 | 1.0112 | 1.0076 |  |  | 0.459
|  |  | 0 | 0
308.9 | 0.466 |  | 0 | 0
329.1 | 0.355 |  | 0 | 0
Alfa_B | 1.192 | Beta_B | 111 | 93
|  |  | 0 | 0
311.5 | 0.45 |  | 0 | 0 |  |  | 5
332.4 | 0.341 |  | 0 | 0
Alfa_C | 0.8265 | Beta_C | 109 | 131
|  |  | 0 | 0
317.5 | 0.414 |  | 0 | 0
339.9 | 0.31 |  | 0 | 0
Alfa_D | 0.36061 | Beta_D | 104 | 278
|  |  | 0 | 0
Alfa_AD | 1.427 | Beta_AD | 10 | 7
Alfa_BD | 0.818 | Beta_BD | 7 | 9
Alfa_CD | 0.46588 | Beta_CD | 5 | 11

## Nfunc083
D83RA | D83NA | D83RC | D83NC | D83RD | D83ND | Rby5 | a | c | d | in pretty print
5 | -1.51 | 5 | 2.18 | 5 | 3.26 | 5 | -1.51 | 2.18 | 3.26 | -1.5 | 2.2 | 3.3 | 0.01 | 0.02 | 0.04
10.74 | 4.7 | 10.76 | 8.39 | 10.72 | 9.62 | 15 | 9.2606 | 12.8971 | 14.1824 | 9.2 | 12.9 | 14.2 | -0.060647 | 0.0029402 | 0.017648
16.48 | 10.83 | 16.51 | 14.49 | 16.44 | 15.7 | 25 | 19.7018 | 23.3471 | 24.6275 | 19.7 | 23.3 | 24.6 | -0.001759 | -0.047139 | -0.027526
22.22 | 16.83 | 22.27 | 20.51 | 22.16 | 21.69 | 35 | 29.9065 | 33.5082 | 34.7616 | 29.9 | 33.5 | 34.8 | -0.0065067 | -0.0081693 | 0.038446
27.96 | 22.74 | 28.02 | 26.46 | 27.88 | 27.58 | 45 | 39.8449 | 43.394 | 44.5867 | 39.8 | 43.4 | 44.6 | -0.044871 | 0.0060056 | 0.013291
33.7 | 28.59 | 33.78 | 32.28 | 33.6 | 33.36 | 55 | 49.3404 | 52.8934 | 54.076 | 49.3 | 52.9 | 54.1 | -0.040402 | 0.0065573 | 0.02402
39.87 | 34.8 | 39.99 | 38.5 | 40.01 | 39.73 | 65 | 58.6907 | 62.1867 | 63.3142 | 58.7 | 62.2 | 63.3 | 0.0093361 | 0.013335 | -0.014152
46.12 | 40.93 | 46.25 | 44.6 | 46.26 | 45.8 | 75 | 67.9965 | 71.4015 | 72.4459 | 68 | 71.4 | 72.4 | 0.0034853 | -0.0015358 | -0.045888
52.43 | 46.93 | 52.58 | 50.62 | 52.56 | 51.79 | 85 | 77.243 | 80.6092 | 81.6241 | 77.2 | 80.6 | 81.6 | -0.043043 | -0.0091746 | -0.024141
58.76 | 52.85 | 58.93 | 56.56 | 58.88 | 57.68 | 95 | 86.5778 | 89.8754 | 90.8361 | 86.6 | 89.9 | 90.8 | 0.022224 | 0.024624 | -0.036083
65.01 | 58.7 | 65.22 | 62.39 | 65.17 | 63.47 | 105 | 95.9718 | 99.1641 | 100.0591 | 96 | 99.2 | 100.1 | 0.028156 | 0.03591 | 0.040888
71.68 | 64.91 | 71.96 | 68.6 | 72.14 | 69.83 | 115 | 105.3164 | 108.4291 | 109.276 | 105.3 | 108.4 | 109.3 | -0.016442 | -0.029128 | 0.024023
78.27 | 71.03 | 78.58 | 74.7 | 78.77 | 75.9 | 125 | 114.7278 | 117.7197 | 118.4987 | 114.7 | 117.7 | 118.5 | -0.027761 | -0.019725 | 0.0013422
84.77 | 77.03 | 85.12 | 80.72 | 85.3 | 81.9 | 135 | 124.2697 | 127.0863 | 127.7279 | 124.3 | 127.1 | 127.7 | 0.030284 | 0.013723 | -0.027928
91.14 | 82.95 | 91.54 | 86.67 | 91.69 | 87.78 | 145 | 133.8086 | 136.5037 | 137.0681 | 133.8 | 136.5 | 137.1 | -0.0086133 | -0.0037106 | 0.031879
97.36 | 88.8 | 97.82 | 92.49 | 97.96 | 93.57 | 155 | 143.4602 | 145.9931 | 146.5049 | 143.5 | 146 | 146.5 | 0.03978 | 0.0069477 | -0.0049043
103.97 | 95.01 | 104.5 | 98.7 | 104.86 | 99.93 | 165 | 153.1836 | 155.5627 | 155.9752 | 153.2 | 155.6 | 156 | 0.016436 | 0.037317 | 0.02477
110.54 | 101.13 | 111.08 | 104.8 | 111.45 | 106.01 | 175 | 162.9608 | 165.2327 | 165.5125 | 163 | 165.2 | 165.5 | 0.039212 | -0.032674 | -0.012494
116.94 | 107.14 | 117.58 | 110.82 | 117.96 | 112 | 185 | 172.7866 | 174.9334 | 175.074 | 172.8 | 174.9 | 175.1 | 0.013394 | -0.033433 | 0.026028
123.23 | 113.05 | 123.98 | 116.77 | 124.34 | 117.89 | 195 | 182.6394 | 184.6057 | 184.6243 | 182.6 | 184.6 | 184.6 | -0.039352 | -0.0057191 | -0.024324
129.38 | 118.9 | 130.22 | 122.59 | 130.61 | 123.67 | 205 | 192.4841 | 194.2836 | 194.1583 | 192.5 | 194.3 | 194.2 | 0.015876 | 0.016413 | 0.041746
135.88 | 125.11 | 136.83 | 128.81 | 137.48 | 130.03 | 215 | 202.4109 | 204.0518 | 203.7711 | 202.4 | 204.1 | 203.8 | -0.010885 | 0.048175 | 0.028851
142.31 | 131.24 | 143.3 | 134.9 | 143.98 | 136.11 | 225 | 212.4107 | 213.9099 | 213.4697 | 212.4 | 213.9 | 213.5 | -0.010729 | -0.0099222 | 0.030315
148.58 | 137.24 | 149.67 | 140.92 | 150.34 | 142.1 | 235 | 222.6125 | 223.9066 | 223.3561 | 222.6 | 223.9 | 223.4 | -0.01248 | -0.0066373 | 0.043946
154.69 | 143.16 | 155.92 | 146.87 | 156.57 | 147.99 | 245 | 232.9094 | 234.043 | 233.3305 | 232.9 | 234 | 233.3 | -0.0094379 | -0.04302 | -0.030472
160.72 | 149 | 162.02 | 152.69 | 162.68 | 153.77 | 255 | 243.0375 | 244.1782 | 243.3484 | 243 | 244.2 | 243.3 | -0.037453 | 0.021839 | -0.048358
167.07 | 155.21 | 168.46 | 158.91 | 169.37 | 160.14 | 265 | 253.2588 | 254.2376 | 253.357 | 253.3 | 254.2 | 253.4 | 0.041246 | -0.037564 | 0.043032
173.34 | 161.34 | 174.77 | 165.01 | 175.73 | 166.21 | 275 | 263.5293 | 264.3569 | 263.3423 | 263.5 | 264.4 | 263.3 | -0.029307 | 0.043099 | -0.042337
179.48 | 167.34 | 180.98 | 171.03 | 181.99 | 172.2 | 285 | 273.6505 | 274.5418 | 273.4295 | 273.7 | 274.5 | 273.4 | 0.049496 | -0.041795 | -0.029511
185.48 | 173.26 | 187.11 | 176.98 | 188.16 | 178.09 |  |  |  |  | 284 | 284.8 | 283.6 | -0.080761 | -0.032459 | 0.10211
191.42 | 179.11 | 193.13 | 182.8 | 194.22 | 183.88
197.72 | 185.32 | 199.56 | 189.01 | 200.89 | 190.24
203.94 | 191.44 | 205.85 | 195.11 | 207.25 | 196.31
210.02 | 197.44 | 212.02 | 201.13 | 213.48 | 202.31
215.95 | 203.36 | 218.08 | 207.08 | 219.6 | 208.2
221.81 | 209.21 | 223.98 | 212.9 | 225.52 | 213.98
227.98 | 215.42 | 230.23 | 219.11 | 231.96 | 220.34
233.97 | 221.55 | 236.29 | 225.21 | 238.08 | 226.42
239.77 | 227.55 | 242.22 | 231.23 | 244.08 | 232.41
245.54 | 233.46 | 248.1 | 237.18 | 249.96 | 238.3
251.31 | 239.31 | 253.83 | 243 | 255.73 | 244.08
257.45 | 245.52 | 260.03 | 249.22 | 262.08 | 250.45
263.44 | 251.65 | 266.07 | 255.32 | 268.18 | 256.52
269.25 | 257.65 | 272.01 | 261.34 | 274.17 | 262.51
275.04 | 263.57 | 277.89 | 267.28 | 280.03 | 268.4
280.82 | 269.42 | 283.59 | 273.11 | 285.75 | 274.19
286.94 | 275.63 | 289.71 | 279.32 | 292.01 | 280.55
292.84 | 281.75 | 295.61 | 285.42 | 298.01 | 286.62
298.49 | 287.75

## comp page
Komyhr adjusted |  |  |  |  | IUP multiply |  |  |  |  | B&P multiply |  |  |  |  | D&M multiply
276.2259 | 283.5894 | 276.34 | 270.2468 | 274.44 | 278.3974 | 294.1189 | 279.1877 | 266.253 | 275.4039 | 276.6963 | 288.6406 | 277.1993 | 267.2536 | 274.2754 | 280.5748 | 296.5796 | 279.4418 | 268.2085 | 275.1278
D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD | D083_TO3_C | D083_TO3_D | D083_TO3_A | D083_TO3_CD | D083_TO3_AD
274.7579 | 279.276 | 275.2198 | 271.1618 | 274.1334 | 276.9179 | 289.6453 | 278.056 | 267.1545 | 275.0962 | 275.2258 | 284.2504 | 276.0757 | 268.1585 | 273.969 | 279.0837 | 292.0686 | 278.3091 | 269.1166 | 274.8204
274.167 | 278.5832 | 275.2163 | 270.6437 | 274.3131 | 276.3223 | 288.9268 | 278.0525 | 266.644 | 275.2766 | 274.6339 | 283.5453 | 276.0721 | 267.6461 | 274.1486 | 278.4835 | 291.3441 | 278.3055 | 268.6023 | 275.0006
275.5552 | 278.828 | 275.2154 | 272.9549 | 274.2507 | 277.7214 | 289.1807 | 278.0515 | 268.9211 | 275.2139 | 276.0245 | 283.7944 | 276.0712 | 269.9317 | 274.0862 | 279.8936 | 291.6001 | 278.3046 | 270.8961 | 274.938
274.5893 | 278.4238 | 275.2952 | 271.5241 | 274.4593 | 276.748 | 288.7614 | 278.1321 | 267.5114 | 275.4233 | 275.0569 | 283.383 | 276.1513 | 268.5167 | 274.2947 | 278.9125 | 291.1773 | 278.3853 | 269.4761 | 275.1471
274.7188 | 279.521 | 275.651 | 270.8585 | 274.6235 | 276.8784 | 289.8994 | 278.4916 | 266.8557 | 275.588 | 275.1866 | 284.4998 | 276.5082 | 267.8585 | 274.4588 | 279.0439 | 292.3249 | 278.7451 | 268.8155 | 275.3117
275.2725 | 280.3489 | 275.9253 | 271.1834 | 274.7549 | 277.4365 | 290.7581 | 278.7688 | 267.1757 | 275.7199 | 275.7412 | 285.3424 | 276.7834 | 268.1798 | 274.5902 | 279.6064 | 293.1907 | 279.0225 | 269.1379 | 275.4435
274.9386 | 279.9295 | 276.1274 | 270.9145 | 275.1209 | 277.1 | 290.3231 | 278.973 | 266.9108 | 276.0872 | 275.4068 | 284.9156 | 276.9861 | 267.9138 | 274.9559 | 279.2673 | 292.7521 | 279.2269 | 268.871 | 275.8104
274.6568 | 280.5399 | 275.7357 | 269.9016 | 274.4687 | 276.816 | 290.9561 | 278.5772 | 265.9129 | 275.4327 | 275.1246 | 285.5368 | 276.5931 | 266.9122 | 274.3041 | 278.981 | 293.3904 | 278.8307 | 267.8658 | 275.1566
275.5259 | 280.0571 | 275.8935 | 271.8685 | 274.795 | 277.6919 | 290.4554 | 278.7367 | 267.8508 | 275.7602 | 275.9951 | 285.0454 | 276.7515 | 268.8573 | 274.6302 | 279.8638 | 292.8855 | 278.9903 | 269.8179 | 275.4837
275.1764 | 282.126 | 276.4339 | 269.5453 | 274.9372 | 277.3396 | 292.6011 | 279.2826 | 265.5619 | 275.9028 | 275.645 | 287.1511 | 277.2935 | 266.5599 | 274.7723 | 279.5088 | 295.0492 | 279.5368 | 267.5122 | 275.6262
275.1516 | 280.8237 | 276.1388 | 270.5587 | 274.906 | 277.3147 | 291.2505 | 278.9844 | 266.5603 | 275.8715 | 275.6202 | 285.8257 | 276.9975 | 267.562 | 274.7411 | 279.4836 | 293.6872 | 279.2383 | 268.518 | 275.595
274.9979 | 282.3793 | 276.0453 | 269.0097 | 274.3826 | 277.1597 | 292.8638 | 278.8899 | 265.0341 | 275.3463 | 275.4661 | 287.4089 | 276.9037 | 266.0301 | 274.218 | 279.3274 | 295.314 | 279.1438 | 266.9806 | 275.0702
275.8524 | 284.4786 | 276.5098 | 268.8477 | 274.4208 | 278.021 | 295.0411 | 279.3592 | 264.8746 | 275.3847 | 276.3221 | 289.5456 | 277.3696 | 265.87 | 274.2563 | 280.1954 | 297.5095 | 279.6135 | 266.8199 | 275.1086
277.2807 | 284.4742 | 276.4652 | 271.4321 | 274.3693 | 279.4605 | 295.0365 | 279.3142 | 267.4208 | 275.333 | 277.7529 | 289.5412 | 277.3249 | 268.4257 | 274.2048 | 281.6462 | 297.5049 | 279.5684 | 269.3848 | 275.057
278.2333 | 282.2703 | 276.7749 | 274.9557 | 275.3358 | 280.4205 | 292.7508 | 279.6271 | 270.8923 | 276.3028 | 278.7071 | 287.298 | 277.6355 | 271.9103 | 275.1707 | 282.6138 | 295.2 | 279.8816 | 272.8818 | 276.0258
276.0617 | 282.8189 | 276.4505 | 270.5669 | 274.7836 | 278.2319 | 293.3197 | 279.2993 | 266.5684 | 275.7487 | 276.5318 | 287.8563 | 277.3101 | 267.5701 | 274.6189 | 280.408 | 295.7738 | 279.5535 | 268.5261 | 275.4723
276.2592 | 282.9149 | 276.1158 | 270.8465 | 274.3367 | 278.431 | 293.4193 | 278.9612 | 266.8438 | 275.3003 | 276.7297 | 287.9541 | 276.9744 | 267.8466 | 274.1722 | 280.6087 | 295.8741 | 279.2151 | 268.8036 | 275.0243
276.1107 | 284.2923 | 276.1387 | 269.4539 | 274.0061 | 278.2813 | 294.8478 | 278.9843 | 265.4719 | 274.9685 | 276.5809 | 289.356 | 276.9974 | 266.4695 | 273.8418 | 280.4578 | 297.3147 | 279.2382 | 267.4215 | 274.6928
275.626 | 285.2853 | 276.2253 | 267.7644 | 273.8561 | 277.7927 | 295.8777 | 279.0718 | 263.8073 | 274.818 | 276.0953 | 290.3667 | 277.0842 | 264.7987 | 273.6919 | 279.9654 | 298.3532 | 279.3258 | 265.7448 | 274.5425
277.0481 | 285.8972 | 276.4006 | 269.8462 | 273.9176 | 279.2261 | 296.5124 | 279.249 | 265.8583 | 274.8797 | 277.5199 | 290.9896 | 277.2601 | 266.8574 | 273.7534 | 281.41 | 298.9931 | 279.5031 | 267.8108 | 274.6042
276.6923 | 284.3115 | 276.6845 | 270.492 | 274.6899 | 278.8675 | 294.8678 | 279.5357 | 266.4946 | 275.6547 | 277.1635 | 289.3756 | 277.5449 | 267.496 | 274.5252 | 281.0486 | 297.3348 | 279.7901 | 268.4518 | 275.3784
276.8015 | 284.1417 | 276.0203 | 270.828 | 273.8968 | 278.9776 | 294.6917 | 278.8647 | 266.8257 | 274.8588 | 277.2729 | 289.2028 | 276.8786 | 267.8284 | 273.7325 | 281.1595 | 297.1572 | 279.1185 | 268.7853 | 274.5832
276.6292 | 283.6013 | 276.5505 | 270.9552 | 274.7068 | 278.8039 | 294.1312 | 279.4004 | 266.9509 | 275.6716 | 277.1003 | 288.6527 | 277.4105 | 267.9541 | 274.542 | 280.9844 | 296.592 | 279.6547 | 268.9114 | 275.3952
276.1764 | 280.7881 | 275.7874 | 272.4252 | 274.4793 | 278.3475 | 291.2136 | 278.6295 | 268.3992 | 275.4433 | 276.6467 | 285.7894 | 276.645 | 269.4078 | 274.3147 | 280.5245 | 293.65 | 278.8831 | 270.3704 | 275.1672
276.2563 | 287.2223 | 276.65 | 267.3271 | 273.8868 | 278.428 | 297.8867 | 279.5009 | 263.3764 | 274.8488 | 276.7267 | 292.3382 | 277.5103 | 264.3662 | 273.7226 | 280.6056 | 300.3789 | 279.7553 | 265.3107 | 274.5733
276.9082 | 285.6775 | 276.5984 | 269.7687 | 274.2254 | 279.0851 | 296.2845 | 279.4488 | 265.7819 | 275.1885 | 277.3798 | 290.7659 | 277.4586 | 266.7807 | 274.0609 | 281.2679 | 298.7633 | 279.7031 | 267.7339 | 274.9127
276.4989 | 285.7632 | 276.7366 | 268.9556 | 274.3774 | 278.6725 | 296.3733 | 279.5884 | 264.9809 | 275.341 | 276.9697 | 290.8531 | 277.5972 | 265.9766 | 274.2128 | 280.8521 | 298.8529 | 279.8429 | 266.9269 | 275.065
277.541 | 287.5598 | 276.6971 | 269.3826 | 273.8584 | 279.7229 | 298.2367 | 279.5485 | 265.4016 | 274.8203 | 278.0136 | 292.6818 | 277.5575 | 266.399 | 273.6942 | 281.9106 | 300.7319 | 279.8029 | 267.3508 | 274.5448
276.6699 | 285.045 | 277.2451 | 269.8506 | 275.2064 | 278.8449 | 295.6285 | 280.1021 | 265.8626 | 276.173 | 277.1411 | 290.1221 | 278.1072 | 266.8617 | 275.0414 | 281.0258 | 298.1019 | 280.3571 | 267.8152 | 275.8962
278.1681 | 284.2484 | 276.8604 | 273.2183 | 274.9294 | 280.3549 | 294.8023 | 279.7135 | 269.1806 | 275.895 | 278.6418 | 289.3113 | 277.7213 | 270.1921 | 274.7645 | 282.5476 | 297.2688 | 279.968 | 271.1575 | 275.6184
276.8549 | 285.2457 | 276.6904 | 270.0222 | 274.4547 | 279.0313 | 295.8366 | 279.5418 | 266.0317 | 275.4186 | 277.3263 | 290.3264 | 277.5509 | 267.0314 | 274.2901 | 281.2137 | 298.3117 | 279.7962 | 267.9855 | 275.1425
277.0886 | 286.0049 | 276.7472 | 269.8275 | 274.3281 | 279.2669 | 296.624 | 279.5991 | 265.84 | 275.2916 | 277.5605 | 291.0991 | 277.6078 | 266.839 | 274.1636 | 281.4511 | 299.1057 | 279.8536 | 267.7923 | 275.0156
277.1325 | 284.4044 | 276.251 | 271.211 | 274.1204 | 279.3111 | 294.9641 | 279.0978 | 267.203 | 275.0831 | 277.6044 | 289.4701 | 277.1101 | 268.2071 | 273.956 | 281.4957 | 297.4319 | 279.3518 | 269.1654 | 274.8074
276.9883 | 282.5283 | 275.9974 | 272.4778 | 274.2905 | 279.1658 | 293.0184 | 278.8415 | 268.451 | 275.2539 | 277.46 | 287.5606 | 276.8556 | 269.4598 | 274.126 | 281.3492 | 295.4699 | 279.0953 | 270.4226 | 274.9779
277.5776 | 291.2055 | 277.6045 | 266.4767 | 274.0512 | 279.7597 | 302.0177 | 280.4652 | 262.5387 | 275.0138 | 278.0503 | 296.3923 | 278.4677 | 263.5253 | 273.8869 | 281.9478 | 304.5445 | 280.7205 | 264.4668 | 274.738
277.0723 | 291.0752 | 277.8768 | 265.6657 | 274.4287 | 279.2505 | 301.8826 | 280.7403 | 261.7396 | 275.3926 | 277.5442 | 296.2597 | 278.7409 | 262.7232 | 274.2641 | 281.4346 | 304.4083 | 280.9958 | 263.6619 | 275.1165
277.3226 | 290.7178 | 277.604 | 266.4108 | 274.178 | 279.5027 | 301.512 | 280.4647 | 262.4737 | 275.141 | 277.7948 | 295.896 | 278.4672 | 263.4601 | 274.0136 | 281.6887 | 304.0346 | 280.72 | 264.4014 | 274.8652